Pipeline Acoustic Emission Dataset

Comprehensive Data Description Report

University of Ulsan | UIAI Lab | BK21 Program Prepared by Muhammad Umar | Supervisor: Prof. Jong-Myon Kim

# Table 1: Experimental Setup

Parameters of the pipeline testbed used for data acquisition. The setup follows the configuration described in Bui Quy & Kim, Measurement 151 (2020) 107150.

| No. | Parameter | Value | Details |
| --- | --- | --- | --- |
| 1 | Pipeline material | Stainless Steel 304 | Industrial-grade, corrosion resistant |
| 2 | Pipeline outer diameter | 114.3 mm | Standard industrial pipe size |
| 3 | Pipeline wall thickness | 6.02 mm | Stainless steel wall |
| 4 | Sensor CH1 position (d1) | 2600 mm | Data file channel 1 = Paper CH1 |
| 5 | Sensor CH2 position (d2) | 100 mm | Data file channel 3 = Paper CH2 |
| 6 | Distance between sensors (L) | 2500 mm | d1 - d2 |
| 7 | Leak position (d) | 900 mm | Fixed position for all experiments |
| 8 | Distance CH1 to leak (l1) | 1700 mm | d1 - d = 2600 - 900 |
| 9 | Distance CH2 to leak (l2) | 800 mm | d - d2 = 900 - 100 |
| 10 | Inner fluid | Water | Water-induced flow |
| 11 | Environmental temperature | ~26 °C | Approximately constant |
| 12 | Wave velocity in water | 1,500 m/s | Used for TDOA computation |
| 13 | Wave velocity in steel | 5,200 m/s | Dominant propagation medium |

# Table 2: AE Sensor Specification

R15I-AST sensors from MISTRAS Corporation used for acoustic emission data acquisition.

| No. | Parameter | Value |
| --- | --- | --- |
| 1 | Sensor model | R15I-AST (MISTRAS) |
| 2 | Peak sensitivity, ref [V/(m/s)] | 109 dB |
| 3 | Peak sensitivity, ref [V/μbar] | -22 dB |
| 4 | Operating frequency range | 50 - 400 kHz |
| 5 | Resonant frequency, ref [V/(m/s)] | 75 kHz |
| 6 | Resonant frequency, ref [V/μbar] | 150 kHz |
| 7 | Directionality | ±1.5 dB |
| 8 | Temperature range | -35 to 75 °C |

# Table 3: Data Acquisition System

| No. | Component | Specification |
| --- | --- | --- |
| 1 | ADC module | NI-9223 (National Instruments) |
| 2 | ADC resolution | 16-bit |
| 3 | Sampling frequency | 1 MHz (1,000,000 Hz) |
| 4 | Number of channels | 3 (Ch1, Ch2, Ch3) |
| 5 | Usable channels | Ch1 and Ch3 only (Ch2 unusable) |
| 6 | Data format | 64-bit double precision (IEEE 754) |
| 7 | Interface | High-speed USB |
| 8 | Recording software | Custom Python + NI library |

# Table 4: .shse File Format Specification

Binary file format used for storing multi-channel AE signal data. Data is stored channel-sequential per second (NOT sample-interleaved).

| Offset (bytes) | Field | Type | Size | Description |
| --- | --- | --- | --- | --- |
| 0 | Timestamp | char[14] | 14 bytes | YYYYMMDDHHmmss format |
| 14 | Sampling rate (fs) | double | 8 bytes | 1,000,000 Hz |
| 22 | Num channels (Nc) | uint8 | 1 byte | 3 channels |
| 23 | Samples per channel | double | 8 bytes | 360,000,000 (6 min × 1 MHz) |
| 31+ | Signal data | double[] | ~8.04 GB | Channel-sequential per second |

Data layout per second: [1,000,000 Ch1 doubles] [1,000,000 Ch2 doubles] [1,000,000 Ch3 doubles]. Total per second = 3,000,000 doubles = 24,000,000 bytes. Total file size = 31 (header) + 360 × 24,000,000 = 8,640,000,031 bytes (~8.04 GB).

# Table 5: Complete Dataset File Inventory

12 .shse files covering 4 leak sizes × 3 pressure levels. Each file contains 6 minutes of 3-channel AE data at 1 MHz sampling rate. Active leak period is approximately 2 minutes within each recording.

| No. | Filename | Leak (mm) | Pressure | Duration | Size (GB) | Timestamp |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 0.3mm_Water_7Bar.shse | 0.3 | 7 Bar | 360 s | 8.04 | 20180821114958 |
| 2 | 0.3mm_Water_13Bar.shse | 0.3 | 13 Bar | 360 s | 8.04 | 20180821114958 |
| 3 | 0.3mm_Water_18Bar.shse | 0.3 | 18 Bar | 360 s | 8.04 | 20180821114958 |
| 4 | 0.5mm_Water_7Bar.shse | 0.5 | 7 Bar | 360 s | 8.04 | 20180821114958 |
| 5 | 0.5mm_Water_13Bar.shse | 0.5 | 13 Bar | 360 s | 8.04 | 20180821114958 |
| 6 | 0.5mm_Water_18Bar.shse | 0.5 | 18 Bar | 360 s | 8.04 | 20180821114958 |
| 7 | 1mm_Water_7Bar.shse | 1.0 | 7 Bar | 360 s | 8.04 | 20180821114958 |
| 8 | 1mm_Water_13Bar.shse | 1.0 | 13 Bar | 360 s | 8.04 | 20180821114958 |
| 9 | 1mm_Water_18Bar.shse | 1.0 | 18 Bar | 360 s | 8.04 | 20180821114958 |
| 10 | 2mm_Water_7Bar.shse | 2.0 | 7 Bar | 360 s | 8.04 | 20180821114958 |
| 11 | 2mm_Water_13Bar.shse | 2.0 | 13 Bar | 360 s | 8.04 | 20180821114958 |
| 12 | 2mm_Water_18Bar.shse | 2.0 | 18 Bar | 360 s | 8.04 | 20180821114958 |

Total dataset size: 12 files × 8.04 GB = 96.5 GB. Total samples: 12 × 360,000,000 × 3 channels = 12,960,000,000 samples.

# Table 6: Channel-to-Sensor Mapping

Critical mapping between data file channels and physical sensor positions. Note: Data file Ch2 is unusable. Per the Description.pptx: 'In the paper, channel 2 is channel 3 in data'.

| Data File Channel | Paper Channel | Position (mm) | Dist. to Leak | Status |
| --- | --- | --- | --- | --- |
| Ch1 (column 1) | CH1 | 2600 | 1700 mm | Active - farther from leak |
| Ch2 (column 2) | N/A | N/A | N/A | UNUSABLE - do not use |
| Ch3 (column 3) | CH2 | 100 | 800 mm | Active - closer to leak |

# Table 7: Signal Statistics per Condition (Corrected Reader, t=60-61s)

Statistics computed using the corrected channel-sequential reader. Ch3 (closer to leak) consistently shows higher RMS than Ch1 (farther), confirming correct channel separation.

| Condition | Ch | Mean | Std | RMS | Peak | Kurtosis | Corr. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2mm_18Bar | 1 | 0.000137 | 0.000473 | 0.000493 | - | - | 0.059 |
| 2mm_18Bar | 3 | 0.000589 | 0.000470 | 0.000754 | - | - | - |
| 2mm_7Bar | 1 | 0.000187 | 0.000443 | 0.000480 | - | - | 0.306 |
| 2mm_7Bar | 3 | 0.000601 | 0.000379 | 0.000710 | - | - | - |
| 0.3mm_18Bar | 1 | 0.000037 | 0.000453 | 0.000455 | - | - | 0.393 |
| 0.3mm_18Bar | 3 | 0.000549 | 0.000461 | 0.000717 | - | - | - |
| 0.3mm_7Bar | 1 | -0.000056 | 0.000411 | 0.000415 | - | - | 0.452 |
| 0.3mm_7Bar | 3 | 0.110226 | 0.000643 | 0.110227 | - | - | - |

Key observation: RMS ratio Ch1/Ch3 ranges from 0.004 to 0.676, confirming Ch3 receives more energy (closer to leak source). Cross-channel correlation is low (0.06-0.45), indicating distinct signal content between channels.

# Table 8: Dominant Frequency Components

Top dominant frequencies identified from Power Spectral Density analysis. The AE signals show a bimodal spectrum: low-frequency flow noise (0.1-0.2 kHz) and high-frequency sensor resonance (379-387 kHz).

| Rank | Frequency (kHz) | Band | Power (V²/Hz) | Source |
| --- | --- | --- | --- | --- |
| 1 | 0.1 | Low (<1 kHz) | 2.64e-11 | Flow turbulence noise |
| 2 | 0.2 | Low (<1 kHz) | 1.53e-11 | Flow turbulence noise |
| 3 | 386.8 | High (300-400 kHz) | 8.80e-12 | Sensor resonance + AE |
| 4 | 379.3 | High (300-400 kHz) | 7.54e-12 | Sensor resonance + AE |
| 5 | 381.5 | High (300-400 kHz) | 7.43e-12 | Sensor resonance + AE |

# Table 9: Dataset Summary at a Glance

| Property | Value |
| --- | --- |
| Total number of files | 12 .shse files |
| Total dataset size | 96.5 GB |
| Total recording duration | 12 × 360s = 72 minutes |
| Total samples (all channels) | 12,960,000,000 |
| Sampling rate | 1 MHz (1,000,000 samples/sec) |
| Number of channels per file | 3 (2 usable: Ch1 and Ch3) |
| Leak sizes tested | 0.3 mm, 0.5 mm, 1.0 mm, 2.0 mm |
| Pressure levels tested | 7 Bar, 13 Bar, 18 Bar |
| Number of conditions | 4 sizes × 3 pressures = 12 |
| Leak position | Fixed at 900 mm |
| Sensor operating band | 50 - 400 kHz |
| ADC resolution | 16-bit |
| Data precision | 64-bit double (float64) |
| File format | .shse (channel-sequential binary) |
| Reference paper | Bui Quy & Kim, Measurement 151 (2020) 107150 |
| Best published result (ABM) | 2.5% mean localization error |

# Table 11: Experimental Conditions Matrix

Each cell shows the filename. F1=2mm, F2=1mm, F3=0.5mm, F4=0.3mm (following the reference paper notation).

| Leak Size | 7 Bar (P1) | 13 Bar (P2) | 18 Bar (P3) |
| --- | --- | --- | --- |
| F4: 0.3 mm | 0.3mm_Water_7Bar | 0.3mm_Water_13Bar | 0.3mm_Water_18Bar |
| F3: 0.5 mm | 0.5mm_Water_7Bar | 0.5mm_Water_13Bar | 0.5mm_Water_18Bar |
| F2: 1.0 mm | 1mm_Water_7Bar | 1mm_Water_13Bar | 1mm_Water_18Bar |
| F1: 2.0 mm | 2mm_Water_7Bar | 2mm_Water_13Bar | 2mm_Water_18Bar |

# Table 12: ABM Reference Localization Results (Bui Quy & Kim, 2019)

Localization results from the reference paper (Table 4). The ABM method achieved 2.5% average error, significantly outperforming CCF (13.3%), GCC-PHAT (16.8%), and other baselines.

| Dataset | CCF | CCF+WDD | CCF+EMD | GCC | ABM | ABM loc | True loc |
| --- | --- | --- | --- | --- | --- | --- | --- |
| F1P1 | 17% | 17% | 15% | 17% | 4% | 800 | 900 |
| F2P1 | 2% | 5% | 1% | 8% | 2% | 850 | 900 |
| F3P1 | 4% | 17% | 17% | 17% | 0% | 900 | 900 |
| F4P1 | 20% | 17% | 17% | 24% | 2% | 950 | 900 |
| F1P2 | 20% | 17% | 17% | 17% | 4% | 800 | 900 |
| F2P2 | 2% | 17% | 3% | 17% | 4% | 800 | 900 |
| F3P2 | 8% | 17% | 1% | 17% | 2% | 850 | 900 |
| F4P2 | 17% | 17% | 17% | 17% | 4% | 800 | 900 |
| F1P3 | 18% | 17% | 21% | 17% | 2% | 850 | 900 |
| F2P3 | 20% | 9% | 11% | 17% | 4% | 800 | 900 |
| F3P3 | 18% | 17% | 7% | 17% | 2% | 850 | 900 |
| F4P3 | 14% | 17% | 19% | 17% | 0% | 900 | 900 |
| Average | 13.3% | 15.3% | 12.2% | 16.8% | 2.5% | 846 | 900 |

# Table 13: Correct Python Code for Reading .shse Files

CRITICAL: The data is stored channel-sequential per second, NOT sample-interleaved. The following Python function correctly reads Ch1 and Ch3.

def read_shse(filepath, start_sec=0, duration_sec=1):     HEADER = 31; nc = 3     with open(filepath, 'rb') as f:         f.read(14)  # timestamp         fs = int(struct.unpack('d', f.read(8))[0])         f.read(1)   # nc         f.read(8)   # nspc         bytes_per_sec = fs * nc * 8         f.seek(HEADER + int(start_sec) * bytes_per_sec)         ch1_list, ch3_list = [], []         for sec in range(int(duration_sec)):             data = np.frombuffer(f.read(bytes_per_sec), dtype=np.float64)             if len(data) < fs * nc: break             ch1_list.append(data[0:fs])       # First fs values             ch3_list.append(data[2*fs:3*fs])  # Last fs values     return np.concatenate(ch1_list), np.concatenate(ch3_list), fs